Biochemistry Uncertainty of Measurement

| No.of copies : | Q-pulse only |
| --- | --- |
| Location of copies : | 1. |
|  | 2. |
|  | 3. |

# Contents

1.	Contents	1

2.	Summary of revisions	1

3.	Introduction	2

4.	Uncertainty of Laboratory Results	2

4.1	Quantitative Tests	2

4.2	Uncertainty of Measurement Data	3

4.3	Calculated Tests	19

4.4	Source of Data	20

5.	References	20

# Summary of revisions

| Version | Summary of change |
| --- | --- |
| 1.0 | New document |
| 2022 | Updated for 2022 |
| 2023 | Under section 4.2 added performance comment and target values to table. |

# Introduction

The Barnsley Biochemistry Laboratory strives to maintain a high standard of performance. Good laboratory practice and effective quality control procedures are invaluable at maintaining this position. However, it should be understood that there is an inherent degree of uncertainty in any laboratory investigation. The level of uncertainty is a combination of three factors; pre-analytical influences, analytical variation and biological variation. This ‘uncertainty’ means that if a series of samples are taken from one individual, for any particular laboratory test, then the results will not be the same.

Although some of the uncertainty can be controlled by the laboratory, it is important for users to recognise the effect of pre-analytical and biological variation when they interpret laboratory tests and to understand the changes in values that must occur before they can be regarded as significant.

# Uncertainty of Laboratory Results

All assays carry an inevitable degree of uncertainty. The Pathology Uncertainty Policy identifies the risks/uncertainties with all the steps in the pre-examination, examination and post examination phase and the control measures that are used to ensure that variation and possibilities for error are minimised. These factors are well recognised but some occur by random error alone. A random error is associated with the fact that when a measurement is repeated it will generally provide a measured value that is different from the previous value. It is random in that the next measured value cannot be predicted exactly from previous such values.

Within the laboratory we monitor analytical imprecision by a variety of methods.

## Quantitative Tests

Analytical variation can be divided into random error (precision) and systematic error (bias). For the monitoring of a patient’s condition, both must be controlled and maintained at an acceptable level. This level may vary from analyte to analyte due to the nature of the assay but can be controlled by the laboratory.

Regular review of External QA samples offers reassurance that accuracy and consistency are being maintained through peer group analysis.

Day to day assurance is provided by repeated analysis of internal quality control material. This analysis occurs at timed intervals. The time period will vary between analysers or specific tests due to manufacturer’s guidelines or the experience – backed up by statistical data – of the laboratory staff. This data is reviewed regularly for trends and required action.

Using a basic interpretation of method mean and Standard Deviation cut off at + and – 2 intervals (see figure 1, below), we are able to access bias and uncertainty for parameters run on these analysers. Additional Westgard statistical rules may also be applied to aid the laboratory to spot biases and trends.

Figure 1 – Typical normal distribution around a mean with cut off points at + and – 2 standard deviations.

As a laboratory generally employs a measurement procedure for long periods of time, the uncertainty of measurement information most relevant to interpreting its test results against fixed reference values is the imprecision of the test results across as many routine operating conditions as possible (for example; multiple calibrator and reagent batches, multiple operators, equipment maintenance, summer/winter etc.). With the caveat that quality control materials may not totally reflect the analytical behaviour of patient specimens, this imprecision is most easily derived from long-term internal quality control (QC) data, calculated as standard deviation (SD) or coefficient of variation (CV%). For the purpose of recording estimates of uncertainty of measurement, the imprecision should be documented as the 95% confidence interval (± 1.96 SD; or ± 1.96 CV%).

## Uncertainty of Measurement Data

Data was collected using the Randox 24/7 statistical programme from 1/01/2021 to 31/03/2021 for all Atellica automated analysis. Please see the data source table below for HbA1c, Ammonia and Osmolality. The average Co-efficient of variation (%CV) was calculated for each analyte.

Please note there are limitations in the data gather. The Randox 24/7 programme contains all IQC measurements and does not exclude those that are beyond the acceptable range used by the laboratory. In these instances, patient samples would not be reported. Consequently, the MoU detailed in the document may overestimate the variation in results. Due to the number of IQC measurements undertaken by the laboratories this effect should be limited.

The range refers to the quoted normal range unless otherwise stated. Where no lower limit is listed this is zero. Please note these values are for serum samples unless stated otherwise.

The following table lists numeric Biochemistry assays performed by the Barnsley Biochemistry Laboratory. The table includes expected target ranges against which laboratory UoM performance can be assessed. Sources of these data are the European Federation of Clinical Chemistry and Laboratory Medicine (EFLM) Biological Variation Database; The Royal College of Pathologists of Australia Quality Assurance Programs (RCPAQAP) Chemical Pathology Analytical Performance Specifications database; or information from the respective manufacturers’ information for users (IFU) documentation when not listed on either independent database.

| Analyte | Reference Range/ Clinical Descision Value |  |  |  | Assay CV (%) | Applied Measurement of Uncertainty |  |  | Target values |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Sex | Age | Category | Concentration |  | Low Concentration | High Concentration | TEa (%) | Target (%) | Source (E=EFLM A=RCPA I=IFU) |
| AFP |  |  | Upper Limit | 10 | 6.79 | 8.67 | 11.33 | 9.52 | 12.00 | A |
| Alcohol* |  |  | Mild intoxication | 50 | 4.16 | 45.92 | 54.08 | 52.0 | 10.00 | A |
|  |  |  | Severe intoxication | 300 | 7.46 | 256.14 | 343.86 | 124.0 | 10.00 | A |
| ALBUMIN |  | Neonate and infant | Lower Limit | 30 | 1.96 | 29 | 31 | 24 | 1.30 | E |
|  |  |  | Upper Limit | 45 | 2.80 | 43 | 47 | 42 | 1.30 | E |
|  |  | 1-16yrs | Lower Limit | 30 | 1.96 | 29 | 31 | 24 | 1.30 | E |
|  |  |  | Upper Limit | 50 | 2.80 | 47 | 53 | 42 | 1.30 | E |
|  |  | Adult | Lower Limit | 35 | 1.96 | 34 | 36 | 24 | 1.30 | E |
|  |  |  | Upper Limit | 50 | 2.80 | 47 | 53 | 42 | 1.30 | E |
| ALKALINE PHOSPHATASE |  | Neonate | Lower Limit | 70 | 5.42 | 63 | 77 | 57 | 2.60 | E |
|  |  |  | Upper Limit | 380 | 4.93 | 343 | 417 | 252 | 2.60 | E |
|  |  | Infant 16 yrs | Lower Limit | 60 | 5.42 | 54 | 66 | 57 | 2.60 | E |
|  |  |  | Upper Limit | 425 | 4.93 | 384 | 466 | 252 | 2.60 | E |
|  |  | > 16 years | Lower Limit | 30 | 5.42 | 27 | 33 | 57 | 2.60 | E |
|  |  |  | Upper Limit | 130 | 4.93 | 117 | 143 | 252 | 2.60 | E |
| ALT |  |  | Lower Limit | 10 | 3.14 | 9 | 11 | 38.8 | 5.00 | E |
|  |  |  | Upper Limit | 40 | 3.14 | 38 | 42 | 38.8 | 5.00 | E |
| Ammonia |  | <28 days | Upper Limit | 100 | 3.30 | 94 | 106 | 168 | 20 | A |
|  |  | >28 days | Upper Limit | 40 | 4.80 | 36 | 44 | 68 | 20 | A |
| AMYLASE |  |  | Lower Limit | 30 | 1.12 | 29.3 | 30.7 | 50.3 | 6.60 | E |
|  |  |  | Upper Limit | 118 | 1.92 | 113.6 | 122.4 | 145 | 6.60 | E |
| AST |  |  | Upper Limit | 34 | 1.91 | 33 | 35 | 41 | 4.80 | E |
| BHCG (Pregnancy) |  |  | Upper Limit | 25 | 6.11 | 22.0 | 28.0 | 19.2 | 10.00 | A |
| BILIRUBIN (TOTAL) |  |  | Upper Limit | 21 | 5.40 | 19 | 23 | 16.9 | 10.00 | E |
| BILIRUBIN (CONJUGATED) |  | <1 month | Lower Limit | 10 | 5.40 | 9 | 11 | 9.55 | 20.00 | A |
|  |  | >1 month | Upper Limit | 7 | 5.40 | 6 | 8 | 9.55 | 20.00 | A |
| BNP |  |  | Lower Limit | 100 | 7.42 | 85 | 115 | 42.7 | 9 | I |
|  |  |  | Upper Limit | 500 | 8.10 | 421 | 579 | 453 | 9 | I |
| C125 |  |  | Upper Limit | 35 | 3.63 | 32.51 | 37.49 | 24.7 | 8.70 | E |
| C199 |  |  | Upper Limit | 35 | 6.70 | 30.40 | 39.60 | 43.9 | 4.30 | E |
| CCP |  |  | Upper Limit | 5 | 7.90 | 4.23 | 5.77 | 3.4 | 8.00 | I |
| CEA* |  |  | Upper Limit | 2.5 | 4.19 | 2.29 | 2.71 | 4.74 | 4.00 | E |
| CALCIUM (Adjusted Serum) |  |  | Lower Limit | 2.2 | 2.00 | 2.11 | 2.29 | 2.13 | 0.90 | E |
|  |  |  | Upper Limit | 2.6 | 2.00 | 2.50 | 2.70 | 2.13 | 0.90 | E |
| CALCIUM (Urine 24 hour) |  |  | Lower Limit | 2.5 | 2.50 | 2.38 | 2.62 | 2.88 | 10.00 | A |
|  |  |  | Upper Limit | 7.5 | 2.50 | 7.13 | 7.87 | 2.88 | 10.00 | A |
| CARBAMAZEPINE (Theraputic Range) |  |  | Lower Limit | 4 | 1.29 | 3.9 | 4.1 | 4.38 | 10.00 | A |
|  |  |  | Upper Limit | 25 | 1.19 | 24.4 | 25.6 | 12.49 | 10.00 | A |
| CHLORIDE (Serum) |  |  | Lower Limit | 98 | 1.37 | 95 | 101 | 99.4 | 0.60 | E |
|  |  |  | Upper Limit | 108 | 1.37 | 105 | 111 | 99.4 | 0.60 | E |
| CHLORIDE (Urine 24 Hour) |  |  | Lower Limit | 110 | 1.03 | 108 | 112 | 164 | 10.00 | A |
|  |  |  | Upper Limit | 250 | 1.03 | 245 | 255 | 164 | 10.00 | S |
| Cholesterol |  |  |  | 5.0 | 2.76 | 4.73 | 5.27 | 5.12 | 2.60 | E |
| CREATINE KINASE |  | Neonates | Upper limit | 900 | 2.80 | 851 | 949 | 511 | 7.50 | E |
|  |  | Infants | Lower Limit | 100 | 9.80 | 81 | 119 | 46.5 | 7.50 | E |
|  |  |  | Upper Limit | 480 | 2.80 | 454 | 506 | 511 | 7.50 | E |
|  | M | Adult | Lower Limit | 40 | 9.80 | 32 | 48 | 46.5 | 7.50 | E |
|  |  |  | Upper Limit | 320 | 2.90 | 302 | 338 | 230 | 7.50 | E |
|  | F | Adult | Lower Limit | 25 | 9.80 | 20 | 30 | 46.5 | 7.50 | E |
|  |  |  | Upper Limit | 200 | 2.90 | 189 | 211 | 230 | 7.50 | E |
| CREATININE (Serum) | M |  | Lower Limit | 53 | 2.50 | 50 | 56 | 58 | 2.30 | E |
|  |  |  | Upper Limit | 97 | 2.50 | 92 | 102 | 58 | 2.30 | E |
|  | F |  | Lower Limit | 44 | 2.50 | 41.8 | 46.2 | 58 | 2.30 | E |
|  |  |  | Upper Limit | 71 | 2.50 | 68 | 74 | 58 | 2.30 | E |
| CREATINE (24 hour urine) |  |  | Lower Limit | 9.0 | 3.00 | 8.5 | 9.5 | 5.9 | 10.00 | A |
|  |  |  | Upper Limit | 17.0 | 3.03 | 16.0 | 18.0 | 13.1 | 10.00 | A |
| CRP |  |  | Upper Limit | 10 | 1.56 | 9.7 | 10.3 | 20 | 29.40 | E |
| CSF GLUCOSE |  |  | Lower Limit | 2.6 | 5.78 | 2.3 | 2.9 | 3.0 | 10.00 | A |
|  |  |  | Upper Limit | 5 | 1.74 | 4.8 | 5.2 | 7.2 | 10.00 | A |
| CSF PROTEIN |  | <1year | Lower Limit | 0.1 | 9.37 | 0.08 | 0.12 | 0.19 | 5.00 | A |
|  |  |  | Upper Limit | 2.1 | 3.86 | 1.9 | 2.3 | 0.85 | 5.00 | A |
|  |  | > 1 year | Lower Limit | 0.1 | 9.37 | 0.08 | 0.12 | 0.19 | 5.00 | A |
|  |  |  | Upper Limit | 2.1 | 3.86 | 1.9 | 2.3 | 0.85 | 5.00 | A |
| CORTISOL (Morning) |  |  | Lower Limit | 140 | 11.50 | 108 | 172 | 237 | 8.20 | E |
|  |  |  | Upper Limit | 500 | 7.70 | 425 | 575 | 792 | 8.20 | E |
| DIGOXIN (Theraputic) |  |  | Lower Limit | 0.5 | 9.80 | 0.4 | 0.6 | 0.78 | 10.00 | A |
|  |  |  | Upper Limit | 2 | 4.90 | 1.8 | 2.2 | 2.5 | 10.00 | A |
| FERRITIN |  |  | Lower Limit | 22 | 10.30 | 18 | 26 | 10 | 6.40 | E |
|  |  |  | Upper Limit | 322 | 4.20 | 295 | 349 | 213 | 6.40 | E |
| FOLATE |  |  | Lower Limit | 3.4 | 9.40 | 2.8 | 4.0 | 6.5 | 5.90 | E |
|  |  |  | Upper Limit | 15.8 | 10.60 | 12.5 | 19.1 | 15.9 | 5.90 | E |
| FREE T3 |  | <3 days | Lower Limit | 2.3 | 3.70 | 2.1 | 2.5 | 4.5 | 2.50 | E |
|  |  |  | Upper Limit | 8.1 | 3.60 | 7.5 | 8.7 | 13.6 | 2.50 | E |
|  |  | <1 month | Lower Limit | 3.5 | 3.70 | 3.2 | 3.8 | 4.5 | 2.50 | E |
|  |  |  | Upper Limit | 6.5 | 3.60 | 6.0 | 7.0 | 13.6 | 2.50 | E |
|  |  | <2 month | Lower Limit | 2.5 | 3.70 | 2.3 | 2.7 | 4.5 | 2.50 | E |
|  |  |  | Upper Limit | 7.8 | 3.60 | 7.2 | 8.4 | 13.6 | 2.50 | E |
|  |  | < 1 year | Lower Limit | 2.7 | 3.70 | 2.5 | 2.9 | 4.5 | 2.50 | E |
|  |  |  | Upper Limit | 7.3 | 3.60 | 6.8 | 7.8 | 13.6 | 2.50 | E |
|  |  | < 5 year | Lower Limit | 3.1 | 3.70 | 2.9 | 3.3 | 4.5 | 2.50 | E |
|  |  |  | Upper Limit | 6.9 | 3.60 | 6.4 | 7.4 | 13.6 | 2.50 | E |
|  |  | <10 year | Lower Limit | 2.3 | 3.70 | 2.1 | 2.5 | 4.5 | 2.50 | E |
|  |  |  | Upper Limit | 6.8 | 3.60 | 6.3 | 7.3 | 13.6 | 2.50 | E |
|  |  | < 14 year | Lower Limit | 3.5 | 3.70 | 3.2 | 3.8 | 4.5 | 2.50 | E |
|  |  |  | Upper Limit | 6.7 | 3.60 | 6.2 | 7.2 | 13.6 | 2.50 | E |
|  |  | < 18 year | Lower Limit | 3.6 | 3.70 | 3.3 | 3.9 | 4.5 | 2.50 | E |
|  |  |  | Upper Limit | 6.7 | 3.60 | 6.2 | 7.2 | 13.6 | 2.50 | E |
|  |  | > 18 year | Lower Limit | 3.5 | 3.70 | 3.2 | 3.8 | 4.5 | 2.50 | E |
|  |  |  | Upper Limit | 6.5 | 3.60 | 6.0 | 7.0 | 13.05 | 2.50 | E |
| FREE T4 |  | <3 days | Lower Limit | 10.8 | 4.53 | 9.8 | 11.8 | 9.6 | 2.50 | E |
|  |  |  | Upper Limit | 26.8 | 2.55 | 25.5 | 28.1 | 22.7 | 2.50 | E |
|  |  | <1 month | Lower Limit | 10.9 | 4.53 | 9.9 | 11.9 | 9.6 | 2.50 | E |
|  |  |  | Upper Limit | 25.5 | 2.55 | 24.2 | 26.8 | 22.7 | 2.50 | E |
|  |  | <2 month | Lower Limit | 11 | 4.53 | 10.0 | 12.0 | 9.6 | 2.50 | E |
|  |  |  | Upper Limit | 24.3 | 2.55 | 23.1 | 25.5 | 22.7 | 2.50 | E |
|  |  | < 1 year | Lower Limit | 11.4 | 4.53 | 10.4 | 12.4 | 9.6 | 2.50 | E |
|  |  |  | Upper Limit | 20.9 | 2.55 | 19.9 | 21.9 | 22.7 | 2.50 | E |
|  |  | < 5 year | Lower Limit | 11.4 | 4.53 | 10.4 | 12.4 | 9.6 | 2.50 | E |
|  |  |  | Upper Limit | 19 | 2.55 | 18.1 | 19.9 | 22.7 | 2.50 | E |
|  |  | <10 year | Lower Limit | 11 | 4.53 | 10.0 | 12.0 | 9.6 | 2.50 | E |
|  |  |  | Upper Limit | 18 | 2.55 | 17.1 | 18.9 | 22.7 | 2.50 | E |
|  |  | < 14 year | Lower Limit | 10.8 | 4.53 | 9.8 | 11.8 | 9.6 | 2.50 | E |
|  |  |  | Upper Limit | 18.7 | 2.55 | 17.8 | 19.6 | 22.7 | 2.50 | E |
|  |  | < 18 year | Lower Limit | 10.7 | 4.53 | 9.7 | 11.7 | 9.6 | 2.50 | E |
|  |  |  | Upper Limit | 18.7 | 2.55 | 17.8 | 19.6 | 22.7 | 2.50 | E |
|  |  | > 18 year | Lower Limit | 10 | 4.53 | 9.1 | 10.9 | 9.6 | 2.50 | E |
|  |  |  | Upper Limit | 18.7 | 2.55 | 17.8 | 19.6 | 22.7 | 2.50 | E |
| FSH (Follicular) | F |  | Lower Limit | 5 | 4.98 | 4.5 | 5.5 | 6.8 | 6.20 | E |
|  |  |  | Upper Limit | 10.2 | 4.98 | 9.2 | 11.2 | 6.8 | 6.20 | E |
| FSH (Mid Cycle) | F |  | Lower Limit | 3.4 | 4.98 | 3.1 | 3.7 | 6.8 | 6.20 | E |
|  |  |  | Upper Limit | 33.4 | 4.23 | 30.6 | 36.2 | 40.5 | 6.20 | E |
| FSH (Luteal) | F |  | Lower Limit | 1.5 | 4.98 | 1.4 | 1.6 | 6.8 | 6.20 | E |
|  |  |  | Upper Limit | 9.1 | 4.98 | 8.2 | 10.0 | 6.8 | 6.20 | E |
| FSH (Post Menopausal) | F |  | Lower Limit | 23 | 3.93 | 21.2 | 24.8 | 6.8 | 6.20 | E |
|  |  |  | Upper Limit | 116.3 | 4.61 | 105.8 | 126.8 | 76.8 | 6.20 | E |
| FSH Children) |  |  | Lower Limit | 0.3 | 4.98 | 0.3 | 0.3 | 6.8 | 6.20 | E |
|  |  |  | Upper Limit | 10.1 | 4.98 | 9.1 | 11.1 | 6.8 | 6.20 | E |
| FASTING GLUCOSE (Serum) |  |  | Indicative of Diabetes Mellitus | 7 | 1.74 | 6.8 | 7.2 | 7.2 | 2.40 | E |
| FASTING GLUCOSE (Whole Blood)** |  |  | Indicative of Diabetes Mellitus | 7 | 1.74 | 6.8 | 7.2 | 7.2 | 2.50 | E |
| Gentamicin |  |  |  | 4 | 3.20 | 3.7 | 4.3 | 4.85 | 6.40 | E |
|  |  |  |  | 6 | 3.20 | 5.6 | 6.4 | 4.85 | 6.40 | E |
| GGT | M |  | Upper Limit | 73 | 2.41 | 69.6 | 76.4 | 67 | 4.50 | E |
|  | F |  | Upper Limit | 38 | 5.50 | 33.9 | 42.1 | 26 | 4.50 | E |
| HbA1c |  |  | Upper Limit | 48 | 1.50 | 46.6 | 49.4 | 33 |  |  |
| HCG (Tumour) |  |  | Upper Limit | 5 | 5.90 | 4.4 | 5.6 | 9.7 | 10.00 | A |
| HDL |  |  |  | 1.00 | 2.25 | 1.0 | 1.0 | 1.03 | 2.90 | E |
| IGA |  | < 13 days | Lower Limit | 0.01 | 2.00 | 0.01 | 0.01 | 1.03 | 2.90 | E |
|  |  |  | Upper Limit | 0.08 | 2.00 | 0.08 | 0.08 | 1.03 | 2.90 | E |
|  |  | 13 - 28 days | Lower Limit | 0.02 | 2.00 | 0.02 | 0.02 | 1.03 | 2.90 | E |
|  |  |  | Upper Limit | 0.15 | 2.00 | 0.14 | 0.16 | 1.03 | 2.90 | E |
|  |  | 28 days – 2 months | Lower Limit | 0.05 | 2.00 | 0.05 | 0.05 | 1.03 | 2.90 | E |
|  |  |  | Upper Limit | 0.4 | 2.00 | 0.38 | 0.42 | 1.03 | 2.90 | E |
|  |  | 2 months – 5 months | Lower Limit | 0.1 | 2.00 | 0.10 | 0.10 | 1.03 | 2.90 | E |
|  |  |  | Upper Limit | 0.5 | 2.00 | 0.48 | 0.52 | 1.03 | 2.90 | E |
|  |  | 5 months – 8 months | Lower Limit | 0.15 | 2.00 | 0.14 | 0.16 | 1.03 | 2.90 | E |
|  |  |  | Upper Limit | 0.7 | 2.00 | 0.67 | 0.73 | 1.03 | 2.90 | E |
|  |  | 8 months – 11 months | Lower Limit | 0.2 | 2.00 | 0.19 | 0.21 | 1.03 | 2.90 | E |
|  |  |  | Upper Limit | 0.7 | 2.00 | 0.67 | 0.73 | 1.03 | 2.90 | E |
|  |  | 11 months – 1 year | Lower Limit | 0.3 | 2.00 | 0.29 | 0.31 | 1.03 | 2.90 | E |
|  |  |  | Upper Limit | 1.2 | 2.00 | 1.15 | 1.25 | 1.03 | 2.90 | E |
|  |  | 1 year – 2 years | Lower Limit | 0.3 | 2.00 | 0.29 | 0.31 | 1.03 | 2.90 | E |
|  |  |  | Upper Limit | 1.3 | 2.00 | 1.25 | 1.35 | 1.03 | 2.90 | E |
|  |  | 2 years – 5 years | Lower Limit | 0.2 | 2.00 | 0.19 | 0.21 | 1.03 | 2.90 | E |
|  |  |  | Upper Limit | 2 | 1.70 | 1.93 | 2.07 | 2.23 | 2.90 | E |
|  |  | 5 years – 8 years | Lower Limit | 0.5 | 2.00 | 0.48 | 0.52 | 1.03 | 2.90 | E |
|  |  |  | Upper Limit | 2.4 | 2.82 | 2.27 | 2.53 | 2.23 | 2.90 | E |
|  |  | 8 years- 11 years | Lower Limit | 0.7 | 1.70 | 0.68 | 0.72 | 1.03 | 2.90 | E |
|  |  |  | Upper Limit | 2.5 | 2.82 | 2.36 | 2.64 | 2.23 | 2.90 | E |
|  |  | 11 years – 44 years | Lower Limit | 0.8 | 1.70 | 0.77 | 0.83 | 1.03 | 2.90 | E |
|  |  |  | Upper Limit | 2.8 | 2.38 | 2.67 | 2.93 | 2.64 | 2.90 | E |
|  |  | > 44 years | Lower Limit | 0.8 | 1.70 | 0.77 | 0.83 | 1.03 | 2.90 | E |
|  |  |  | Upper Limit | 4 | 2.38 | 3.81 | 4.19 | 2.64 | 2.90 | E |
| IGG |  | < 13 days | Lower Limit | 5 | 3.10 | 4.70 | 5.30 | 7.1 | 1.80 | E |
|  |  |  | Upper Limit | 17 | 2.60 | 16.13 | 17.87 | 16.9 | 1.80 | E |
|  |  | 13 - 28 days | Lower Limit | 3.9 | 3.10 | 3.66 | 4.14 | 7.1 | 1.80 | E |
|  |  |  | Upper Limit | 13 | 2.70 | 12.31 | 13.69 | 12.2 | 1.80 | E |
|  |  | 28 days – 2 months | Lower Limit | 2.1 | 3.10 | 1.97 | 2.23 | 7.1 | 1.80 | E |
|  |  |  | Upper Limit | 7.7 | 3.10 | 7.23 | 8.17 | 7.1 | 1.80 | E |
|  |  | 2 months – 5 months | Lower Limit | 2.4 | 3.10 | 2.25 | 2.55 | 7.1 | 1.80 | E |
|  |  |  | Upper Limit | 8.8 | 3.10 | 8.27 | 9.33 | 7.1 | 1.80 | E |
|  |  | 5 months – 8 months | Lower Limit | 3 | 3.10 | 2.82 | 3.18 | 7.1 | 1.80 | E |
|  |  |  | Upper Limit | 9 | 3.10 | 8.45 | 9.55 | 7.1 | 1.80 | E |
|  |  | 8 months – 11 months | Lower Limit | 3 | 3.10 | 2.82 | 3.18 | 7.1 | 1.80 | E |
|  |  |  | Upper Limit | 10.9 | 2.70 | 10.32 | 11.48 | 12.2 | 1.80 | E |
|  |  | 11 months – 1 year | Lower Limit | 3.1 | 3.10 | 2.91 | 3.29 | 7.1 | 1.80 | E |
|  |  |  | Upper Limit | 13.8 | 2.70 | 13.07 | 14.53 | 12.2 | 1.80 | E |
|  |  | 1 year – 2 years | Lower Limit | 3.7 | 3.10 | 3.48 | 3.92 | 7.1 | 1.80 | E |
|  |  |  | Upper Limit | 15.8 | 2.60 | 14.99 | 16.61 | 16.9 | 1.80 | E |
|  |  | 2 years – 5 years | Lower Limit | 4.9 | 3.10 | 4.60 | 5.20 | 7.1 | 1.80 | E |
|  |  |  | Upper Limit | 16.1 | 2.60 | 15.28 | 16.92 | 16.9 | 1.80 | E |
|  |  | 5 years – 14 years | Lower Limit | 5.4 | 3.10 | 5.07 | 5.73 | 7.1 | 1.80 | E |
|  |  |  | Upper Limit | 16.1 | 2.60 | 15.28 | 16.92 | 16.9 | 1.80 | E |
|  |  | > 14 years | Lower Limit | 6 | 3.10 | 5.64 | 6.36 | 7.1 | 1.80 | E |
|  |  |  | Upper Limit | 16 | 2.60 | 15.18 | 16.82 | 16.9 | 1.80 | E |
| IGM |  | < 13 days | Lower Limit | 0.05 | 4.70 | 0.05 | 0.05 | 0.41 | 3.00 | E |
|  |  |  | Upper Limit | 2 | 3.00 | 1.88 | 2.12 | 1.36 | 3.00 | E |
|  |  | 13 - 28 days | Lower Limit | 0.08 | 4.70 | 0.07 | 0.09 | 0.41 | 3.00 | E |
|  |  |  | Upper Limit | 0.4 | 4.70 | 0.36 | 0.44 | 0.41 | 3.00 | E |
|  |  | 28 days – 2 months | Lower Limit | 0.17 | 4.70 | 0.15 | 0.19 | 0.41 | 3.00 | E |
|  |  |  | Upper Limit | 0.7 | 4.70 | 0.64 | 0.76 | 0.41 | 3.00 | E |
|  |  | 2 months – 5 months | Lower Limit | 0.2 | 4.70 | 0.18 | 0.22 | 0.41 | 3.00 | E |
|  |  |  | Upper Limit | 1 | 4.70 | 0.91 | 1.09 | 0.41 | 3.00 | E |
|  |  | 5 months – 8 months | Lower Limit | 0.4 | 4.70 | 0.36 | 0.44 | 0.41 | 3.00 | E |
|  |  |  | Upper Limit | 1.6 | 3.00 | 1.51 | 1.69 | 1.36 | 3.00 | E |
|  |  | 8 months – 11 months | Lower Limit | 0.6 | 4.70 | 0.54 | 0.66 | 0.41 | 3.00 | E |
|  |  |  | Upper Limit | 2.1 | 1.17 | 2.05 | 2.15 | 1.36 | 3.00 | E |
|  |  | 11 months – 2 year | Lower Limit | 0.5 | 4.70 | 0.45 | 0.55 | 0.41 | 3.00 | E |
|  |  |  | Upper Limit | 2.2 | 3.00 | 2.07 | 2.33 | 1.36 | 3.00 | E |
|  |  | 2 year – 5 years | Lower Limit | 0.5 | 4.70 | 0.45 | 0.55 | 0.41 | 3.00 | E |
|  |  |  | Upper Limit | 2 | 3.00 | 1.88 | 2.12 | 1.36 | 3.00 | E |
|  |  | 5 years – 11 years | Lower Limit | 0.5 | 4.70 | 0.45 | 0.55 | 0.41 | 3.00 | E |
|  |  |  | Upper Limit | 1.8 | 3.00 | 1.69 | 1.91 | 1.36 | 3.00 | E |
|  |  | 11 years – 44 years | Lower Limit | 0.5 | 4.70 | 0.45 | 0.55 | 0.41 | 3.00 | E |
|  |  |  | Upper Limit | 1.9 | 3.00 | 1.79 | 2.01 | 1.36 | 3.00 | E |
| IRON | M |  | Lower Limit | 12 | 3.70 | 11 | 13 | 15.6 | 10.30 | E |
|  |  |  | Upper Limit | 31 | 2.30 | 30 | 32 | 26 | 10.30 | E |
|  | F |  | Lower Limit | 9 | 3.70 | 8 | 10 | 15.6 | 10.30 | E |
|  |  |  | Upper Limit | 30 | 2.30 | 29 | 31 | 26 | 10.30 | E |
| LACTATE DEHYDROGENASE |  |  | Lower Limit | 120 | 2.93 | 113 | 127 | 145.0 | 2.60 | E |
|  |  |  | Upper Limit | 246 | 2.60 | 233 | 259 | 241.0 | 2.60 | E |
| LH (Follicular levels) | F |  | Lower Limit | 1.9 | 3.31 | 1.8 | 2.0 | 2.42 | 11.40 | E |
|  |  |  | Upper Limit | 12.5 | 3.31 | 11.7 | 13.3 | 2.42 | 11.40 | E |
| LH (Peak levels) | F |  | Lower Limit | 8.7 | 3.31 | 8.1 | 9.3 | 2.42 | 11.40 | E |
|  |  |  | Upper Limit | 76.3 | 4.00 | 70.3 | 82.3 | 33.6 | 11.40 | E |
| LH (Luteal) | F |  | Lower Limit | 0.5 | 3.31 | 0.5 | 0.5 | 2.42 | 11.40 | E |
|  |  |  | Upper Limit | 16.9 | 3.31 | 15.8 | 18.0 | 2.42 | 11.40 | E |
| LH (Post Menopausall) | F |  | Lower Limit | 7.9 | 3.31 | 7 | 8 | 2.42 | 11.40 | E |
|  |  |  | Upper Limit | 53.8 | 4.00 | 50 | 58 | 33.6 | 11.40 | E |
| LH | M | <70yrs | Lower Limit | 1.5 | 3.31 | 1.4 | 1.6 | 2.42 | 11.40 | E |
|  |  |  | Upper Limit | 9.3 | 3.31 | 8.7 | 9.9 | 2.42 | 11.40 | E |
| LITHIUM (Theraputic Range) |  |  | Lower Limit | 0.4 | 3.60 | 0.4 | 0.4 | 0.91 | 10.00 | A |
|  |  |  | Upper Limit | 1 | 3.60 | 0.9 | 1.1 | 0.91 | 10.00 | A |
| MAGNESIUM (Blood) |  | Neonate | Lower Limit | 0.6 | 5.60 | 0.5 | 0.7 | 0.47 | 10.00 | A |
|  |  |  | Upper Limit | 1 | 3.30 | 0.9 | 1.1 | 0.93 | 10.00 | A |
|  |  | Infant -16yrs | Lower Limit | 0.7 | 5.60 | 0.6 | 0.8 | 0.47 | 10.00 | A |
|  |  |  | Upper Limit | 1 | 3.30 | 0.9 | 1.1 | 0.93 | 2.00 | E |
|  |  | Adult | Lower Limit | 0.7 | 5.60 | 0.6 | 0.8 | 0.47 | 2.00 | E |
|  |  |  | Upper Limit | 1 | 3.30 | 0.9 | 1.1 | 0.93 | 2.00 | E |
| MAGNESIUM (24hr Urine) |  |  | Lower Limit | 2.4 | 4.17 | 2.2 | 2.6 | 2.05 | 2.00 | E |
|  |  |  | Upper Limit | 6.5 | 3.60 | 6.0 | 7.0 | 4.3 | 2.00 | E |
| Microalbumin (Urine) |  |  |  | 10 | 15.10 | 7.0 | 13.0 | 7.1 | 2.00 | E |
|  |  |  |  | 65 | 2.70 | 61.6 | 68.4 | 57.6 | 2.00 | E |
| Oestradiol |  | Male 15ys | Upper Limit | 44 | 11.50 | 34.1 | 53.9 | 151 | 2.00 | E |
|  |  | Male Adult | Upper Limit | 150 | 11.50 | 116.2 | 183.8 | 151 | 10.00 | A |
| Osmolality |  |  | Lower Limit | 275 | 1.10 | 269.1 | 280.9 | 305 | 3.00 | A |
|  |  |  | Upper Limit | 295 | 1.10 | 288.6 | 301.4 | 305 | 3.00 | A |
| Osmolality (Urine |  |  | Lower Limit | 50 | 1.50 | 48.5 | 51.5 | 439 | 2 | A |
|  |  |  | Upper Limit | 1200 | 1.10 | 1174.1 | 1225.9 | 800 | 2 | A |
| Paracetamol (serum) |  |  |  | 150 | 2.49 | 142.7 | 157.3 | 164 | 10.00 | A |
| PHENYTOIN (Theraputic Range) |  |  | Lower Limit | 5 | 4.10 | 4.6 | 5.4 | 6.5 | 10.00 | A |
|  |  |  | Upper Limit | 20 | 3.70 | 18.5 | 21.5 | 19.5 | 10.00 | A |
| POTASSIUM (Serum) |  | Neonate | Lower Limit | 3.4 | 0.82 | 3.3 | 3.5 | 2.6 | 2.00 | E |
|  |  |  | Upper Limit | 6.0 | 1.70 | 5.8 | 6.2 | 5.8 | 2.00 | E |
|  |  | Infant | Lower Limit | 3.5 | 1.90 | 3.4 | 3.6 | 3.8 | 2.00 | E |
|  |  |  | Upper Limit | 5.7 | 1.70 | 5.5 | 5.9 | 5.8 | 2.00 | E |
|  |  | 1-16yrs | Lower Limit | 2.5 | 0.82 | 2.5 | 2.5 | 2.6 | 2.00 | E |
|  |  |  | Upper Limit | 6.5 | 1.70 | 6.3 | 6.7 | 5.8 | 2.00 | E |
|  |  | Adult | Lower Limit | 3.5 | 1.90 | 3.4 | 3.6 | 3.8 | 2.00 | E |
|  |  |  | Upper Limit | 5.3 | 1.70 | 5.1 | 5.5 | 5.8 | 2.00 | E |
| POTASSIUM (Urine) |  |  | Lower Limit | 40 | 2.10 | 38.4 | 41.6 | 30.4 | 10.00 | A |
|  |  |  | Upper Limit | 120 | 0.90 | 117.9 | 122.1 | 76 | 10.00 | A |
| PHOSPHATE (Serum) |  | < 1 month | Lower Limit | 1.4 | 2.67 | 1.33 | 1.47 | 1.23 | 15.00 | A |
|  |  |  | Upper Limit | 3.1 | 2.86 | 2.93 | 3.27 | 2.22 | 3.90 | E |
|  |  | 1 month - 1 year | Lower Limit | 1 | 2.50 | 0.95 | 1.05 | 1.13 | 3.90 | E |
|  |  |  | Upper Limit | 2.6 | 2.86 | 2.45 | 2.75 | 2.22 | 3.90 | E |
|  |  | 1 year- 16 years | Lower Limit | 1.2 | 2.50 | 1.14 | 1.26 | 1.13 | 3.90 | E |
|  |  |  | Upper Limit | 2.3 | 2.86 | 2.17 | 2.43 | 2.22 | 3.90 | E |
|  |  | > 16 years | Lower Limit | 0.8 | 2.50 | 0.76 | 0.84 | 1.13 | 3.90 | E |
|  |  |  | Upper Limit | 1.5 | 2.50 | 1.43 | 1.57 | 1.13 | 3.90 | E |
| PHOSPHATE (Urine) |  |  | Lower Limit | 15 | 5.30 | 13.4 | 16.6 | 14.7 | 15.00 | A |
|  |  |  | Upper Limit | 50 | 5.30 | 44.8 | 55.2 | 14.7 | 15.00 | A |
| PROGESTERONE (Follicular levels) | F |  | Upper Limit | 4.45 | 10.30 | 3.6 | 5.3 | 2.5 | 9.30 | E |
| PROGESTERONE (Luteal) | F |  | Lower Limit | 10.3 | 7.78 | 8.7 | 11.9 | 2.5 | 9.30 | E |
|  |  |  | Upper Limit | 81.28 | 5.30 | 72.8 | 89.7 | 78.4 | 9.30 | E |
| PROGESTERONE (Mid Luteal) | F |  | Lower Limit | 14.12 | 7.78 | 12.0 | 16.3 | 25.90 | 9.30 | E |
|  |  |  | Upper Limit | 89.14 | 5.30 | 79.9 | 98.4 | 78.4 | 9.30 | E |
| PROGESTERONE | F (Post Menopausal) |  | Upper Limit | 2.32 | 10.30 | 1.9 | 2.8 | 2.5 | 9.30 | E |
|  | M |  | Upper Limit | 388 | 5.30 | 347.7 | 428.3 | 78.4 | 9.30 | E |
| PROLACTIN | F |  | Upper Limit | 646 | 3.91 | 596.5 | 695.5 | 745 | 14.00 | E |
| PROTEIN (Serum) |  |  | Lower Limit | 60 | 2.90 | 57 | 63 | 67.6 | 1.30 | E |
|  |  |  | Upper Limit | 80 | 3.00 | 75 | 85 | 92 | 1.30 | E |
| PROTEIN (24 Hour Urine) |  |  | Upper Limit | 0.15 | 9.40 | 0.12 | 0.18 | 0.19 | 10.00 | A |
| PROTEIN (URINE) |  |  | Upper Limit | 0.1 | 9.40 | 0.08 | 0.12 | 0.19 | 10.00 | A |
| PSA |  |  | Upper Limit | 3 | 4.10 | 2.76 | 3.24 | 1.58 | 6.80 | E |
| PTH |  |  | Lower Limit | 1.95 | 3.00 | 1.84 | 2.06 | 3.5 | 14.70 | E |
|  |  |  | Upper Limit | 8.49 | 3.00 | 7.99 | 8.99 | 3.5 | 14.70 | E |
| RF |  |  | Upper Limit | 13 | 14.30 | 9.36 | 16.64 | 11.3 |  |  |
| SALICYLATE |  |  | Theraputic upper limit | 250 | 4.40 | 228 | 272 | 140 | 10.00 | A |
|  |  |  | Severe toxicity | 500 | 2.52 | 475 | 525 | 512 | 10.00 | A |
| SHBG | M | <50 yrs | Lower Limit | 14.6 | 7.30 | 13 | 17 | 11 | 9.70 | E |
|  |  |  | Upper Limit | 94.6 | 5.10 | 85 | 104 | 41.7 | 9.70 | E |
|  | M | >50 yrs | Lower Limit | 21.6 | 7.30 | 19 | 25 | 11 | 9.70 | E |
|  |  |  | Upper Limit | 113.3 | 5.10 | 102 | 125 | 41.7 | 9.70 | E |
| SODIUM (Serum) |  |  | Lower Limit | 133 | 1.13 | 130 | 136 | 138 | 0.30 | E |
|  |  |  | Upper Limit | 146 | 1.13 | 143 | 149 | 138 | 0.30 | E |
| SODIUM (Urine 24 hour) |  |  | Lower Limit | 40 | 1.04 | 39 | 41 | 82 | 10.00 | A |
|  |  |  | Upper Limit | 220 | 0.80 | 217 | 223 | 165 | 10.00 | A |
| TESTOSTERONE | M |  | Lower Limit | 9.3 | 4.90 | 8 | 10 | 14.6 | 12.50 | E |
|  |  |  | Upper Limit | 32.2 | 4.90 | 29 | 35 | 14.6 | 12.50 | E |
| THEOPHYLLINE (Theraputic Range) |  |  | 12.9 | 12 | 4.98 | 10.8 | 13.2 | 13.1 | 10.00 | A |
|  |  |  | Upper Limit | 20 | 4.98 | 18.0 | 22.0 | 13.1 | 10.00 | A |
| TROPONIN I |  |  | Detectable cardiac damage | 3 | 7.10 | 3 | 3 | 43 | 6.00 | E |
|  |  |  | MI Rule In | 120 | 6.40 | 105 | 135 | 125.4 | 6.00 | E |
| TRANSFERRIN |  |  | Lower Limit | 2 | 2.60 | 1.9 | 2.1 | 2.64 | 1.90 | E |
|  |  |  | Upper Limit | 3.2 | 2.60 | 3.0 | 3.4 | 2.64 | 1.90 | E |
| TSH |  | < 3 days | Lower Limit | 0.13 | 3.90 | 0.12 | 0.14 | 0.07 | 8.80 | E |
|  |  |  | Upper Limit | 9.23 | 4.70 | 8.38 | 10.08 | 21.2 | 8.80 | E |
|  |  | < 1 month | Lower Limit | 0.16 | 3.90 | 0.15 | 0.17 | 0.07 | 8.80 | E |
|  |  |  | Upper Limit | 8.48 | 4.70 | 7.70 | 9.26 | 21.2 | 8.80 | E |
|  |  | <2 month | Lower Limit | 0.19 | 3.90 | 0.18 | 0.20 | 0.07 | 8.80 | E |
|  |  |  | Upper Limit | 7.78 | 4.70 | 7.06 | 8.50 | 21.2 | 8.80 | E |
|  |  | < 1 year | Lower Limit | 0.3 | 3.90 | 0.28 | 0.32 | 0.07 | 8.80 | E |
|  |  |  | Upper Limit | 5.88 | 4.70 | 5.34 | 6.42 | 21.2 | 8.80 | E |
|  |  | < 5 year | Lower Limit | 0.42 | 3.90 | 0.39 | 0.45 | 0.07 | 8.80 | E |
|  |  |  | Upper Limit | 4.79 | 4.70 | 4.35 | 5.23 | 21.2 | 8.80 | E |
|  |  | <10 year | Lower Limit | 0.48 | 3.90 | 0.44 | 0.52 | 0.07 | 8.80 | E |
|  |  |  | Upper Limit | 4.67 | 4.70 | 4.24 | 5.10 | 21.2 | 8.80 | E |
|  |  | < 14 year | Lower Limit | 0.53 | 3.90 | 0.49 | 0.57 | 0.07 | 8.80 | E |
|  |  |  | Upper Limit | 4.58 | 4.70 | 4.16 | 5.00 | 21.2 | 8.80 | E |
|  |  | < 18 year | Lower Limit | 0.56 | 3.90 | 0.52 | 0.60 | 0.07 | 8.80 | E |
|  |  |  | Upper Limit | 4.53 | 4.70 | 4.11 | 4.95 | 21.2 | 8.80 | E |
|  |  | > 18 year | Lower Limit | 0.38 | 3.90 | 0.35 | 0.41 | 0.07 | 8.80 | E |
|  |  |  | Upper Limit | 5.5 | 4.50 | 5.01 | 5.99 | 2.48 | 8.80 | E |
| UREA |  | < 1 month | Lower Limit | 0.8 | 7.50 | 0.7 | 0.9 | 2.9 | 7.00 | E |
|  |  |  | Upper Limit | 5.5 | 3.18 | 5.2 | 5.8 | 9.8 | 7.00 | E |
|  |  | < 1 year | Lower Limit | 1 | 7.50 | 0.9 | 1.1 | 2.9 | 7.00 | E |
|  |  |  | Upper Limit | 5.1 | 3.18 | 4.8 | 5.4 | 9.8 | 7.00 | E |
|  |  | < 16 year | Lower Limit | 2.5 | 7.50 | 2.1 | 2.9 | 2.9 | 7.00 | E |
|  |  |  | Upper Limit | 6.5 | 3.18 | 6.1 | 6.9 | 9.8 | 7.00 | E |
|  |  | > 16 years | Lower Limit | 2.5 | 7.50 | 2.1 | 2.9 | 2.9 | 7.00 | E |
|  |  |  | Upper Limit | 7.8 | 3.18 | 7.3 | 8.3 | 9.8 | 7.00 | E |
| Urea (Urine) |  | 24 Hour | Lower Limit | 250 | 3.10 | 234.8 | 265.2 | 267 | 10.00 | A |
|  |  |  | Upper Limit | 600 | 3.10 | 563.5 | 636.5 | 267 | 10.00 | A |
| URIC ACID | M |  | Lower Limit | 200 | 1.54 | 194 | 206 | 250 | 4.20 | E |
|  |  |  | Upper Limit | 430 | 2.29 | 411 | 449 | 501 | 4.20 | E |
|  | F |  | Lower Limit | 140 | 1.54 | 136 | 144 | 250 | 4.20 | E |
|  |  |  | Upper Limit | 360 | 2.29 | 344 | 376 | 501 | 4.20 | E |
| Urin Acid (Urine) |  | 24 Hour | Lower Limit | 1.5 | 1.60 | 1.5 | 1.5 | 0.9 | 15.00 | A |
|  |  |  | Upper Limit | 4.5 | 1.60 | 4.4 | 4.6 | 0.9 | 15.00 | A |
| VALPROATE (Theraputic Range) |  |  | Lower Limit | 50 | 3.30 | 47 | 53 | 41.7 | 10.00 | A |
|  |  |  | Upper Limit | 100 | 3.44 | 93 | 107 | 80.4 | 10.00 | A |
| Vitamin B12 |  |  | Lower Limit | 211 | 9.80 | 170.47 | 251.53 | 211 | 15.00 | A |
|  |  |  | Upper Limit | 911 | 5.90 | 805.65 | 1016.35 | 926 | 15.00 | A |
| Vitamin D |  |  | Deficiency | 25 | 11.70 | 19.3 | 30.7 | 19.3 | 6.90 | E |
|  |  |  | Adequate | 50 | 9.80 | 40.4 | 59.6 | 66.2 | 6.90 | E |
|  |  |  | Toxic | 220 | 7.30 | 188.5 | 251.5 | 130.9 | 6.90 | E |

* Data from original spreadsheet recalculated with fliers removed

The data for Blood Gas MoU was obtained from the EQA precision data (attached)

## Calculated Tests

## Source of Data

The following QC materials were used to establish this performance data:

Barnsley Biochemistry

| IQC Used | Lot | Date Range | Analyser |  |
| --- | --- | --- | --- | --- |
| Randox Liquid Chemistry 1 | 367 | 01/10/2022-18/11/2022 | Atellica 1 and 2 |  |
| Randox Liquid Chemistry 2 | 1582 | 01/10/2022-18/11/2022 | Atellica 1 and 2 |  |
| Randox Liquid Chemistry 3 | 1256 | 01/10/2022-18/11/2022 | Atellica 1 and 2 |  |
| Randox Urine 2 | 1160 | 01/10/2022-18/11/2022 | Atellica 1 and 2 |  |
| Randox Urine 3 | 1161 | 01/10/2022-18/11/2022 | Atellica 1 and 2 |  |
| Randox Immunoassay 1 | 2105 | 01/10/2022-18/11/2022 | Atellica 1 and 2 |  |
| Randox Immunoassay 2 | 2124 | 01/10/2022-18/11/2022 | Atellica 1 and 2 |  |
| Randox Immunoassay 3 | 2079 | 01/10/2022-18/11/2022 | Atellica 1 and 2 |  |
| BioRad Cardiac 1B | 67675 | 01/10/2022-18/11/2022 | Atellica 1 and 2 |  |
| BioRad Cardiac 1C | 67671 | 01/10/2022-18/11/2022 | Atellica 1 and 2 |  |
| BioRad Cardiac 2 | 67672 | 01/10/2022-18/11/2022 | Atellica 1 and 2 |  |
| Siemens PTH | 3910621 | 01/10/2022-18/11/2022 | Atellica 1 |  |
| Siemens PTH | 3910622 | 01/10/2022-18/11/2022 | Atellica 1 |  |
| Siemens PTH | 3910623 | 01/10/2022-18/11/2022 | Atellica 1 |  |
| Siemens aCCP | 604531 | 01/10/2022-18/11/2022 | Atellica 2 |  |
| Siemens aCCP | 604532 | 01/10/2022-18/11/2022 | Atellica 2 |  |
| Quantametrix Microalbumin 1 | 41301 | 01/10/2022-18/11/2022 | Atellica 1 and 2 |  |
| Tosoh HbA1c | 3010 | 01/10/2022-18/11/2022 | Tosoh 1 and 2 |  |
| Humatrol | 1389 | 01/09/2022-18/11/2022 | Osmometer |  |
| Randox Urine 2 | 1160 | 01/09/2022-18/11/2022 | Osmometer |  |
| Randox Urine 3 | 1161 | 01/09/2022-18/11/2022 | Osmometer |  |
| Blood Gas EQA | Internal CVP1 and CVP2 | 01/05/2022 – 01/11/2022 | GEM 5000 |  |

# References

G.H. White, I. Farrance

“Uncertainty of Measurement in Quantitative Medical Testing- A Users Guide”

Clin Biochem Rev Vol 25 Suppl (ii) November 2004 I S1